## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.2743 | 72.1 | 0.1425 | 0.27425 | 2.7425 | 0.58 | 18.2967 | 96.4 | 0.1898 | 8.2967 | 82.9672 | 12.47 | 41.36
20 | 20.8278 | 102.6 | 0.203 | 0.8278 | 4.139 | 0.61 | 27.8947 | 118.6 | 0.2352 | 7.8947 | 39.4736 | 20.42 | 58.72
30 | 30.5662 | 123.8 | 0.2469 | 0.56622 | 1.8874 | 0.74 | 38.9063 | 139.8 | 0.2783 | 8.9063 | 29.6878 | 32.79 | 73.36
40 | 41.8464 | 145.3 | 0.288 | 1.8464 | 4.616 | 0.82 | 49.0458 | 157.4 | 0.3116 | 9.0458 | 22.6146 | 46.21 | 85.61
50 | 50.7517 | 160.1 | 0.317 | 0.7517 | 1.5034 | 0.83 | 60.4275 | 175 | 0.3453 | 10.4275 | 20.855 | 57.87 | 94.15
60 | 61.8623 | 176.8 | 0.3499 | 1.8623 | 3.1039 | 0.87 | 70.2351 | 188.5 | 0.3726 | 10.2351 | 17.0585 | 70.95 | 2.92
70 | 70.2744 | 188 | 0.3738 | 0.2744 | 0.392 | 0.89 | 81.4826 | 203.3 | 0.4008 | 11.4826 | 16.4038 | 82.96 | 10.45
80 | 83.004 | 205.1 | 0.4047 | 3.004 | 3.755 | 0.9 | 91.2908 | 214.6 | 0.4254 | 11.2908 | 14.1136 | 94.83 | 17.62
90 | 91.1411 | 214.5 | 0.4249 | 1.1411 | 1.2678 | 0.88 | 104.8047 | 231 | 0.4537 | 14.8047 | 16.4497 | 104.22 | 22.09
100 | 103.7399 | 228.1 | 0.4548 | 3.7399 | 3.7399 | 0.89 | 115.219 | 241.6 | 0.4769 | 15.219 | 15.219 | 116.9 | 27.33

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage