## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.3891 | 72.6 | 0.1431 | 0.38906 | 3.8906 | 0.77 | 11.9273 | 77.4 | 0.1541 | 1.9273 | 19.2734 | 15.45 | 30.21
20 | 20.0997 | 99.8 | 0.2014 | 0.09972 | 0.4986 | 0.86 | 23.3279 | 108.2 | 0.2156 | 3.3279 | 16.6396 | 28.9 | 44.61
30 | 30.6652 | 123.6 | 0.2481 | 0.66516 | 2.2172 | 0.95 | 32.7803 | 128.5 | 0.2551 | 2.7803 | 9.2678 | 43.01 | 55.34
40 | 40.356 | 142.5 | 0.2832 | 0.356 | 0.89 | 0.92 | 45.0296 | 150.5 | 0.2992 | 5.0296 | 12.574 | 53.76 | 60.74
50 | 51.7934 | 161.3 | 0.3211 | 1.7934 | 3.5869 | 0.94 | 55.4443 | 166.8 | 0.3324 | 5.4443 | 10.8886 | 66.39 | 65.9
60 | 62.5514 | 177.3 | 0.3528 | 2.5514 | 4.2524 | 0.95 | 66.0108 | 182.3 | 0.3621 | 6.0108 | 10.0181 | 79.1 | 70.31
70 | 73.1913 | 191.5 | 0.3822 | 3.1913 | 4.559 | 0.95 | 76.3814 | 195.9 | 0.3899 | 6.3814 | 9.1163 | 91.19 | 74
80 | 83.6198 | 204.8 | 0.4083 | 3.6198 | 4.5248 | 0.95 | 87.4656 | 209.8 | 0.4169 | 7.4656 | 9.332 | 102.73 | 77.74
90 | 94.2428 | 217.5 | 0.4333 | 4.2428 | 4.7142 | 0.98 | 96.3812 | 220.3 | 0.4375 | 6.3813 | 7.0903 | 115.74 | 81.98
100 | 104.2402 | 229.2 | 0.4548 | 4.2402 | 4.2402 | 0.94 | 108.4134 | 233.7 | 0.4639 | 8.4134 | 8.4134 | 122.98 | 83.79
110 | 115.4872 | 241 | 0.4792 | 5.4872 | 4.9884 | 0.98 | 116.8371 | 242.3 | 0.4822 | 6.8371 | 6.2155 | 140.23 | 88.01
120 | 125.4995 | 251.2 | 0.4996 | 5.4995 | 4.5829 | 0.99 | 126.1256 | 252.1 | 0.5003 | 6.1256 | 5.1047 | 153.31 | 91.16
130 | 135.067 | 261.1 | 0.5173 | 5.067 | 3.8977 | 1 | 135.4322 | 261.2 | 0.5185 | 5.4322 | 4.1786 | 165.41 | 93.87
140 | 145.9831 | 271.9 | 0.5369 | 5.9831 | 4.2737 | 1 | 146.2294 | 271.6 | 0.5384 | 6.2294 | 4.4496 | 178.09 | 96.24
150 | 156.6018 | 280.9 | 0.5575 | 6.6017 | 4.4012 | 1 | 157.4124 | 282 | 0.5582 | 7.4124 | 4.9416 | 191.16 | 98.49

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage