## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.0534 | 71.2 | 0.1412 | 0.05344 | 0.5344 | 0.69 | 13.9278 | 83.5 | 0.1668 | 3.9278 | 39.278 | 13.38 | 33.86
20 | 20.9202 | 102.5 | 0.2041 | 0.92025 | 4.6013 | 0.81 | 25.0318 | 112.2 | 0.2231 | 5.0318 | 25.1591 | 26.08 | 51.51
30 | 30.8986 | 125.4 | 0.2464 | 0.89856 | 2.9952 | 0.88 | 34.2557 | 131.5 | 0.2605 | 4.2557 | 14.1858 | 38.99 | 61.97
40 | 41.9049 | 145.1 | 0.2888 | 1.9049 | 4.7622 | 0.89 | 47.0018 | 153.4 | 0.3064 | 7.0018 | 17.5044 | 50.92 | 69.06
50 | 51.0401 | 159.7 | 0.3196 | 1.0401 | 2.0802 | 0.85 | 58.344 | 171.6 | 0.34 | 8.344 | 16.688 | 60.14 | 73.53
60 | 62.4797 | 177.6 | 0.3518 | 2.4797 | 4.1328 | 0.92 | 66.7206 | 183.5 | 0.3636 | 6.7206 | 11.201 | 75.61 | 81.59
70 | 73.4009 | 192.3 | 0.3817 | 3.4009 | 4.8584 | 0.95 | 76.6343 | 196.7 | 0.3896 | 6.6343 | 9.4776 | 89.57 | 87.72
80 | 83.5787 | 204.9 | 0.4079 | 3.5787 | 4.4734 | 0.94 | 88.7669 | 211.4 | 0.4199 | 8.7669 | 10.9586 | 99.98 | 91.7
90 | 93.5911 | 217.3 | 0.4307 | 3.5911 | 3.9901 | 0.95 | 98.9206 | 223.6 | 0.4424 | 8.9206 | 9.9118 | 112.23 | 96.19
100 | 104.7385 | 230.6 | 0.4542 | 4.7385 | 4.7385 | 0.94 | 109.9052 | 236 | 0.4657 | 9.9052 | 9.9052 | 122.63 | 99.38
110 | 115.3386 | 241.8 | 0.477 | 5.3386 | 4.8533 | 0.96 | 119.2634 | 245.6 | 0.4856 | 9.2634 | 8.4212 | 137.45 | 3.24
120 | 125.4714 | 251.9 | 0.4981 | 5.4714 | 4.5595 | 0.97 | 129.3586 | 255.7 | 0.5059 | 9.3586 | 7.7989 | 149.78 | 6.4
130 | 136.3435 | 262.3 | 0.5198 | 6.3435 | 4.8796 | 0.98 | 139.6241 | 265.9 | 0.5251 | 9.6241 | 7.4031 | 163 | 9.51
140 | 146.4974 | 272.3 | 0.538 | 6.4974 | 4.641 | 0.98 | 149.2995 | 274.7 | 0.5435 | 9.2994 | 6.6425 | 175.3 | 13.8
150 | 155.9246 | 281.3 | 0.5543 | 5.9246 | 3.9497 | 0.97 | 160.9941 | 285.4 | 0.5641 | 10.9941 | 7.3294 | 185.5 | 17.22

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage