## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.4907 | 72.3 | 0.1451 | 0.49073 | 4.9073 | 0.67 | 14.3227 | 84.8 | 0.1689 | 4.3227 | 43.2272 | 16.38 | 25.21
20 | 20.574 | 101.6 | 0.2025 | 0.574 | 2.87 | 0.96 | 20.02 | 100.1 | 0.2 | 0.02 | 0.1 | 28.8 | 41.96
30 | 31.2499 | 125.2 | 0.2496 | 1.2499 | 4.1664 | 0.97 | 31.9033 | 126.5 | 0.2522 | 1.9033 | 6.3443 | 38.98 | 52.79
40 | 41.4288 | 144.1 | 0.2875 | 1.4287 | 3.5719 | 0.98 | 42.1515 | 145.3 | 0.2901 | 2.1515 | 5.3788 | 52.05 | 59.95
50 | 52.0191 | 161 | 0.3231 | 2.0191 | 4.0382 | 0.96 | 54.2686 | 164.9 | 0.3291 | 4.2686 | 8.5372 | 63.17 | 65.17
60 | 62.5518 | 176.9 | 0.3536 | 2.5518 | 4.2531 | 0.95 | 65.305 | 180.7 | 0.3614 | 5.305 | 8.8416 | 73.45 | 69.86
70 | 73.2838 | 192.7 | 0.3803 | 3.2838 | 4.6912 | 0.99 | 73.1911 | 191.7 | 0.3818 | 3.1911 | 4.5587 | 88.44 | 75.95
80 | 82.7193 | 204.7 | 0.4041 | 2.7193 | 3.3991 | 0.96 | 86.7986 | 208.3 | 0.4167 | 6.7986 | 8.4983 | 97.58 | 79.36
90 | 93.8956 | 217.2 | 0.4323 | 3.8956 | 4.3284 | 0.95 | 98.7231 | 222.6 | 0.4435 | 8.7231 | 9.6923 | 107.25 | 82.81
100 | 100.5534 | 224.6 | 0.4477 | 0.55342 | 0.55342 | 0.95 | 111.8141 | 237.7 | 0.4704 | 11.8141 | 11.8141 | 118.91 | 86.84
110 | 114.2655 | 239.5 | 0.4771 | 4.2655 | 3.8777 | 0.98 | 118.6231 | 245.8 | 0.4826 | 8.6231 | 7.8392 | 134.55 | 91.21
120 | 123.9518 | 249.4 | 0.497 | 3.9518 | 3.2932 | 0.96 | 131.3022 | 256.6 | 0.5117 | 11.3022 | 9.4185 | 143.52 | 93.67
130 | 134.6025 | 259.7 | 0.5183 | 4.6025 | 3.5404 | 0.96 | 137.9942 | 263.7 | 0.5233 | 7.9942 | 6.1494 | 155.24 | 97.07
140 | 146.3403 | 270.8 | 0.5404 | 6.3403 | 4.5288 | 0.99 | 147.0982 | 271.8 | 0.5412 | 7.0982 | 5.0701 | 171.25 | 0.88
150 | 154.6235 | 278.3 | 0.5556 | 4.6235 | 3.0823 | 0.97 | 160.7432 | 284.3 | 0.5654 | 10.7432 | 7.1621 | 179.52 | 2.62
160 | 166.868 | 289.4 | 0.5766 | 6.868 | 4.2925 | 0.98 | 171.4617 | 293.8 | 0.5836 | 11.4617 | 7.1635 | 192.02 | 5.11
170 | 175.9428 | 297 | 0.5924 | 5.9428 | 3.4958 | 1 | 175.112 | 296.8 | 0.59 | 5.112 | 3.0071 | 207.32 | 8.06
180 | 188.344 | 307.4 | 0.6127 | 8.344 | 4.6355 | 1 | 186.048 | 306 | 0.608 | 6.048 | 3.36 | 218.72 | 11.38
190 | 197.6292 | 315.5 | 0.6264 | 7.6292 | 4.0154 | 1 | 198.7273 | 317 | 0.6269 | 8.7273 | 4.5933 | 230.51 | 14.72
200 | 206.2694 | 325.5 | 0.6337 | 6.2694 | 3.1347 | 1 | 202.4703 | 318.6 | 0.6355 | 2.4703 | 1.2351 | 241.2 | 17.99

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage