## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.498 | 72.3 | 0.1452 | 0.49796 | 4.9796 | 0.67 | 14.792 | 86.1 | 0.1718 | 4.792 | 47.9198 | 16.35 | 25.17
20 | 20.0901 | 100.3 | 0.2003 | 0.09009 | 0.45045 | 1 | 20.1001 | 100.4 | 0.2002 | 0.10008 | 0.5004 | 29.5 | 42.92
30 | 31.1875 | 125 | 0.2495 | 1.1875 | 3.9583 | 0.98 | 31.8908 | 126.3 | 0.2525 | 1.8907 | 6.3025 | 39.03 | 52.83
40 | 41.4863 | 144.2 | 0.2877 | 1.4863 | 3.7159 | 0.98 | 42.2678 | 145.4 | 0.2907 | 2.2678 | 5.6695 | 51.95 | 59.91
50 | 52.325 | 162.5 | 0.322 | 2.325 | 4.65 | 0.96 | 54.384 | 164.9 | 0.3298 | 4.384 | 8.768 | 63.26 | 65.21
60 | 62.4988 | 176.9 | 0.3533 | 2.4988 | 4.1646 | 0.94 | 65.4134 | 181 | 0.3614 | 5.4134 | 9.0223 | 73.36 | 69.83
70 | 72.9046 | 191.2 | 0.3813 | 2.9046 | 4.1494 | 0.99 | 73.3826 | 191.9 | 0.3824 | 3.3826 | 4.8322 | 88.27 | 75.89
80 | 82.2958 | 203.3 | 0.4048 | 2.2958 | 2.8698 | 0.97 | 86.7567 | 208.6 | 0.4159 | 6.7567 | 8.4459 | 98.04 | 79.52
90 | 93.5492 | 216.8 | 0.4315 | 3.5492 | 3.9436 | 0.95 | 98.5232 | 222.4 | 0.443 | 8.5232 | 9.4702 | 107.51 | 82.9
100 | 100.4416 | 224.1 | 0.4482 | 0.44162 | 0.44162 | 0.94 | 110.2374 | 234 | 0.4711 | 10.2374 | 10.2374 | 117.06 | 86.2
110 | 113.2165 | 238.1 | 0.4755 | 3.2165 | 2.9241 | 0.97 | 118.0973 | 243.6 | 0.4848 | 8.0973 | 7.3612 | 133.58 | 90.97
120 | 123.8518 | 249.5 | 0.4964 | 3.8518 | 3.2098 | 0.96 | 130.5344 | 256 | 0.5099 | 10.5344 | 8.7787 | 143.44 | 93.65
130 | 134.914 | 260 | 0.5189 | 4.914 | 3.78 | 1 | 136.4506 | 261.6 | 0.5216 | 6.4506 | 4.962 | 160.13 | 98.48
140 | 146.1489 | 271.4 | 0.5385 | 6.1489 | 4.3921 | 0.99 | 146.3405 | 270.6 | 0.5408 | 6.3405 | 4.5289 | 171.19 | 0.86
150 | 156.24 | 280 | 0.558 | 6.24 | 4.16 | 0.98 | 161.3677 | 284.8 | 0.5666 | 11.3677 | 7.5785 | 179.62 | 2.65
160 | 165.6564 | 288.6 | 0.574 | 5.6564 | 3.5352 | 0.97 | 171.9606 | 294 | 0.5849 | 11.9606 | 7.4754 | 190.36 | 4.79
170 | 176.1771 | 298.1 | 0.591 | 6.1771 | 3.6336 | 1 | 175.9121 | 297.4 | 0.5915 | 5.9121 | 3.4777 | 207.32 | 8.06
180 | 188.4956 | 308.1 | 0.6118 | 8.4956 | 4.7198 | 1 | 188.4052 | 307.5 | 0.6127 | 8.4053 | 4.6696 | 218.72 | 11.38
190 | 197.6631 | 314.7 | 0.6281 | 7.6631 | 4.0332 | 1 | 199.143 | 316.1 | 0.63 | 9.143 | 4.8121 | 230.17 | 14.62
200 | 202.7566 | 318.9 | 0.6358 | 2.7566 | 1.3783 | 0.99 | 207.807 | 319.9 | 0.6496 | 7.807 | 3.9035 | 238.74 | 17.24

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.1517 | 72.1 | 0.1408 | 0.15168 | 1.5168 | 0.76 | 207.807 | 319.9 | 0.6496 | 7.807 | 3.9035 | 15.32 | 30.06
20 | 20.1603 | 100.4 | 0.2008 | 0.16032 | 0.8016 | 0.85 | 23.7947 | 109.2 | 0.2179 | 3.7947 | 18.9734 | 28.49 | 44.2
30 | 30.7147 | 124.2 | 0.2473 | 0.71466 | 2.3822 | 0.92 | 33.7739 | 130.2 | 0.2594 | 3.7739 | 12.5796 | 42.1 | 54.8
40 | 40.5552 | 142.8 | 0.284 | 0.5552 | 1.388 | 0.91 | 45.6473 | 151.1 | 0.3021 | 5.6473 | 14.1183 | 53.01 | 60.4
50 | 52.067 | 161.9 | 0.3216 | 2.067 | 4.1341 | 0.93 | 56.2464 | 168 | 0.3348 | 6.2464 | 12.4928 | 66.03 | 65.76
60 | 62.8345 | 177.8 | 0.3534 | 2.8345 | 4.7242 | 0.94 | 66.1018 | 182.4 | 0.3624 | 6.1018 | 10.1696 | 78.52 | 70.12
70 | 73.2096 | 192 | 0.3813 | 3.2096 | 4.5851 | 0.95 | 77.047 | 196.9 | 0.3913 | 7.047 | 10.0671 | 90.73 | 73.85
80 | 83.4154 | 204.6 | 0.4077 | 3.4154 | 4.2693 | 0.95 | 87.1316 | 209.3 | 0.4163 | 7.1316 | 8.9145 | 102.03 | 77.53
90 | 93.5918 | 217.1 | 0.4311 | 3.5918 | 3.9909 | 0.98 | 96.558 | 220 | 0.4389 | 6.558 | 7.2867 | 115.52 | 81.9
100 | 100.7778 | 224.8 | 0.4483 | 0.77784 | 0.77784 | 0.93 | 108.624 | 233.6 | 0.465 | 8.624 | 8.624 | 122.18 | 83.61
110 | 115.1494 | 241.1 | 0.4776 | 5.1494 | 4.6812 | 0.98 | 116.8358 | 242.7 | 0.4814 | 6.8358 | 6.2143 | 140.18 | 88
120 | 125.8 | 251.6 | 0.5 | 5.8 | 4.8333 | 1 | 125.3732 | 251.4 | 0.4987 | 5.3732 | 4.4776 | 154.09 | 91.33
130 | 135.3785 | 261.5 | 0.5177 | 5.3786 | 4.1373 | 1 | 134.9656 | 260.3 | 0.5185 | 4.9656 | 3.8197 | 165.41 | 93.87
140 | 146.2 | 272 | 0.5375 | 6.2 | 4.4286 | 1 | 145.8785 | 271.2 | 0.5379 | 5.8785 | 4.1989 | 178.09 | 96.24
150 | 156.7398 | 281.4 | 0.557 | 6.7398 | 4.4932 | 1 | 156.8518 | 281.5 | 0.5572 | 6.8518 | 4.5679 | 191.16 | 98.49

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage