## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.1173 | 71.5 | 0.1415 | 0.11725 | 1.1725 | 0.69 | 14.003 | 83.8 | 0.1671 | 4.003 | 40.0298 | 13.43 | 33.93
20 | 20.4322 | 101.3 | 0.2017 | 0.43221 | 2.161 | 0.79 | 27.3428 | 116.7 | 0.2343 | 7.3428 | 36.7141 | 25.56 | 51.03
30 | 31.2875 | 125.2 | 0.2499 | 1.2875 | 4.2916 | 0.87 | 34.6896 | 132 | 0.2628 | 4.6896 | 15.632 | 38.86 | 61.88
40 | 41.9328 | 145.6 | 0.288 | 1.9328 | 4.832 | 0.88 | 46.9557 | 153.4 | 0.3061 | 6.9557 | 17.3894 | 50.38 | 68.78
50 | 52.0187 | 161.8 | 0.3215 | 2.0187 | 4.0374 | 0.87 | 59.1678 | 172.3 | 0.3434 | 9.1678 | 18.3356 | 61.04 | 74.02
60 | 62.8525 | 177.7 | 0.3537 | 2.8525 | 4.7541 | 0.92 | 67.2149 | 184.1 | 0.3651 | 7.2149 | 12.0249 | 75.08 | 81.33
70 | 72.4656 | 190.9 | 0.3796 | 2.4656 | 3.5223 | 0.94 | 77.4005 | 197.4 | 0.3921 | 7.4005 | 10.5722 | 88.58 | 87.32
80 | 83.947 | 205.4 | 0.4087 | 3.947 | 4.9337 | 0.93 | 89.9476 | 213.5 | 0.4213 | 9.9476 | 12.4344 | 98.97 | 91.33
90 | 92.9448 | 216 | 0.4303 | 2.9448 | 3.272 | 0.93 | 99.9669 | 225.1 | 0.4441 | 9.9669 | 11.0743 | 110.71 | 95.63
100 | 102.8731 | 228.1 | 0.451 | 2.8731 | 2.8731 | 0.93 | 111.153 | 237 | 0.469 | 11.153 | 11.153 | 121.13 | 99
110 | 115.366 | 241.2 | 0.4783 | 5.366 | 4.8781 | 0.95 | 120.1214 | 246.1 | 0.4881 | 10.1214 | 9.2013 | 135.82 | 2.78
120 | 125.8725 | 252.3 | 0.4989 | 5.8725 | 4.8937 | 0.96 | 130.941 | 257.1 | 0.5093 | 10.941 | 9.1175 | 149.3 | 6.29
130 | 136.0812 | 262.3 | 0.5188 | 6.0812 | 4.6779 | 0.97 | 139.8389 | 265.5 | 0.5267 | 9.8388 | 7.5683 | 161.52 | 9.09
140 | 146.6901 | 271.9 | 0.5395 | 6.69 | 4.7786 | 0.97 | 149.6825 | 275 | 0.5443 | 9.6825 | 6.9161 | 174.1 | 13.38
150 | 157.0231 | 280.9 | 0.559 | 7.0231 | 4.6821 | 0.97 | 161.9904 | 286 | 0.5664 | 11.9904 | 7.9936 | 184.34 | 16.84

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage