## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.1387 | 70.9 | 0.143 | 0.1387 | 1.387 | 0.65 | 14.0114 | 83.7 | 0.1674 | 4.0114 | 40.1138 | 16.17 | 24.87
20 | 20.6552 | 101.8 | 0.2029 | 0.65522 | 3.2761 | 0.97 | 20.04 | 100.2 | 0.2 | 0.04 | 0.2 | 28.85 | 42.02
30 | 31.3501 | 125.1 | 0.2506 | 1.3501 | 4.5002 | 0.97 | 31.9792 | 126.4 | 0.253 | 1.9792 | 6.5973 | 38.78 | 52.65
40 | 41.6305 | 144.3 | 0.2885 | 1.6305 | 4.0764 | 0.99 | 42.2823 | 145.4 | 0.2908 | 2.2823 | 5.7058 | 52.47 | 60.16
50 | 51.8579 | 160.7 | 0.3227 | 1.8579 | 3.7158 | 0.96 | 54.4664 | 164.8 | 0.3305 | 4.4664 | 8.9328 | 63.16 | 65.16
60 | 61.6882 | 175.5 | 0.3515 | 1.6883 | 2.8137 | 0.95 | 65.54 | 181.2 | 0.3617 | 5.54 | 9.2334 | 73.7 | 69.98
70 | 73.134 | 191.3 | 0.3823 | 3.134 | 4.4771 | 0.99 | 73.4402 | 191.8 | 0.3829 | 3.4402 | 4.9146 | 88.08 | 75.81
80 | 83.62 | 204.7 | 0.4085 | 3.6199 | 4.5249 | 0.96 | 86.9027 | 208.7 | 0.4164 | 6.9027 | 8.6283 | 97.84 | 79.45
90 | 93.9394 | 216.8 | 0.4333 | 3.9394 | 4.3772 | 0.95 | 98.5679 | 222.2 | 0.4436 | 8.5679 | 9.5199 | 107.76 | 82.99
100 | 101.248 | 226 | 0.448 | 1.248 | 1.248 | 0.94 | 104.9278 | 229.1 | 0.458 | 4.9278 | 4.9278 | 116.85 | 86.13
110 | 115.0319 | 240.1 | 0.4791 | 5.0319 | 4.5745 | 0.98 | 117.4666 | 242.8 | 0.4838 | 7.4666 | 6.7879 | 134.79 | 91.27
120 | 125.3753 | 250.4 | 0.5007 | 5.3753 | 4.4794 | 0.96 | 130.9952 | 256 | 0.5117 | 10.9952 | 9.1627 | 143.28 | 93.61
130 | 134.5244 | 258.9 | 0.5196 | 4.5244 | 3.4803 | 0.98 | 135.9763 | 262.3 | 0.5184 | 5.9763 | 4.5972 | 157.12 | 97.62
140 | 146.4485 | 270.9 | 0.5406 | 6.4485 | 4.6061 | 1 | 147.7227 | 272.3 | 0.5425 | 7.7227 | 5.5162 | 171.5 | 0.93
150 | 156.856 | 280.1 | 0.56 | 6.856 | 4.5707 | 0.98 | 161.0848 | 284 | 0.5672 | 11.0848 | 7.3899 | 180.67 | 2.86
160 | 165.8303 | 288.2 | 0.5754 | 5.8303 | 3.6439 | 0.97 | 172.0782 | 294 | 0.5853 | 12.0782 | 7.5489 | 190.75 | 4.86
170 | 175.6466 | 296.5 | 0.5924 | 5.6466 | 3.3215 | 1 | 176.4164 | 297.9 | 0.5922 | 6.4164 | 3.7743 | 207.32 | 8.06
180 | 186.5576 | 307.8 | 0.6061 | 6.5576 | 3.6431 | 1 | 186.2931 | 305.9 | 0.609 | 6.2931 | 3.4962 | 218.72 | 11.38
190 | 198.4729 | 317.1 | 0.6259 | 8.4729 | 4.4594 | 0.99 | 199.1102 | 316.5 | 0.6291 | 9.1102 | 4.7948 | 228.96 | 14.28
200 | 202.9464 | 319.5 | 0.6352 | 2.9464 | 1.4732 | 1 | 206.2704 | 319.7 | 0.6452 | 6.2704 | 3.1352 | 241.2 | 17.99

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.0747 | 71.3 | 0.1413 | 0.07469 | 0.7469 | 0.69 | 13.9946 | 83.6 | 0.1674 | 3.9946 | 39.9464 | 13.41 | 33.91
20 | 20.1402 | 100.4 | 0.2006 | 0.14024 | 0.7012 | 0.78 | 25.8397 | 112.2 | 0.2303 | 5.8397 | 29.1983 | 25.43 | 50.92
30 | 31.1749 | 125.1 | 0.2492 | 1.1749 | 3.9164 | 0.87 | 34.348 | 131.4 | 0.2614 | 4.348 | 14.4932 | 38.94 | 61.94
40 | 41.5294 | 144.4 | 0.2876 | 1.5294 | 3.8236 | 0.88 | 47.0778 | 154 | 0.3057 | 7.0778 | 17.6945 | 50.27 | 68.72
50 | 51.232 | 160.1 | 0.32 | 1.232 | 2.464 | 0.86 | 59.0298 | 172.4 | 0.3424 | 9.0298 | 18.0595 | 60.41 | 73.67
60 | 62.6222 | 177.4 | 0.353 | 2.6222 | 4.3703 | 0.92 | 66.7947 | 183.2 | 0.3646 | 6.7947 | 11.3245 | 75.25 | 81.41
70 | 72.615 | 191.9 | 0.3784 | 2.615 | 3.7357 | 0.95 | 76.9679 | 197 | 0.3907 | 6.9679 | 9.9541 | 88.82 | 87.42
80 | 83.5167 | 205.1 | 0.4072 | 3.5167 | 4.3959 | 0.93 | 88.5784 | 210.5 | 0.4208 | 8.5784 | 10.723 | 99.56 | 91.54
90 | 93.6132 | 217.2 | 0.431 | 3.6132 | 4.0147 | 0.94 | 99.9894 | 225.1 | 0.4442 | 9.9894 | 11.0994 | 111.56 | 95.95
100 | 104.4901 | 229.8 | 0.4547 | 4.4901 | 4.4901 | 0.94 | 109.9545 | 235.6 | 0.4667 | 9.9545 | 9.9545 | 123.12 | 99.5
110 | 115.437 | 241.5 | 0.478 | 5.437 | 4.9427 | 0.95 | 119.8037 | 245.6 | 0.4878 | 9.8037 | 8.9124 | 136.32 | 2.92
120 | 125.7968 | 252.3 | 0.4986 | 5.7968 | 4.8307 | 0.97 | 130.2773 | 256.3 | 0.5083 | 10.2773 | 8.5644 | 149.59 | 6.36
130 | 136.2138 | 262 | 0.5199 | 6.2138 | 4.7798 | 0.97 | 139.9375 | 266.7 | 0.5247 | 9.9375 | 7.6442 | 162.02 | 9.2
140 | 146.7707 | 272.1 | 0.5394 | 6.7707 | 4.8362 | 0.98 | 149.3544 | 274.7 | 0.5437 | 9.3544 | 6.6817 | 174.24 | 13.43
150 | 154.3144 | 279.1 | 0.5529 | 4.3144 | 2.8763 | 0.97 | 161.1374 | 285.3 | 0.5648 | 11.1374 | 7.425 | 184.32 | 16.84

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage