## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.1102 | 70.8 | 0.1428 | 0.11024 | 1.1024 | 0.66 | 14.3566 | 84.8 | 0.1693 | 4.3566 | 43.5664 | 16.27 | 25.03
20 | 20.1101 | 100.4 | 0.2003 | 0.11012 | 0.5506 | 0.95 | 20.7974 | 101.5 | 0.2049 | 0.79735 | 3.9868 | 28.58 | 41.65
30 | 31.3376 | 125.1 | 0.2505 | 1.3376 | 4.4585 | 0.98 | 31.9539 | 126.3 | 0.253 | 1.9539 | 6.513 | 39.3 | 53.04
40 | 40.8122 | 142.9 | 0.2856 | 0.81224 | 2.0306 | 0.97 | 42.2677 | 145.5 | 0.2905 | 2.2677 | 5.6694 | 51.59 | 59.73
50 | 51.9063 | 160.9 | 0.3226 | 1.9063 | 3.8127 | 0.96 | 54.483 | 164.9 | 0.3304 | 4.483 | 8.9659 | 63.18 | 65.17
60 | 62.5695 | 176.8 | 0.3539 | 2.5695 | 4.2825 | 0.94 | 64.3687 | 179.4 | 0.3588 | 4.3687 | 7.2812 | 73.39 | 69.84
70 | 73.3828 | 191.6 | 0.383 | 3.3828 | 4.8326 | 0.99 | 73.5357 | 192.2 | 0.3826 | 3.5357 | 5.051 | 88.71 | 76.05
80 | 83.4155 | 204.5 | 0.4079 | 3.4156 | 4.2694 | 0.97 | 86.861 | 208.6 | 0.4164 | 6.861 | 8.5763 | 98.04 | 79.52
90 | 93.9394 | 216.8 | 0.4333 | 3.9394 | 4.3772 | 0.95 | 98.7679 | 222.3 | 0.4443 | 8.7679 | 9.7421 | 107.66 | 82.96
100 | 102.0164 | 226 | 0.4514 | 2.0164 | 2.0164 | 0.94 | 110.9911 | 235.8 | 0.4707 | 10.9911 | 10.9911 | 116.95 | 86.17
110 | 114.9361 | 239.9 | 0.4791 | 4.9361 | 4.4874 | 0.97 | 118.6332 | 243.7 | 0.4868 | 8.6332 | 7.8483 | 133.24 | 90.88
120 | 123.4794 | 248.8 | 0.4963 | 3.4794 | 2.8995 | 0.95 | 130.9952 | 256 | 0.5117 | 10.9952 | 9.1627 | 143.19 | 93.58
130 | 136.1107 | 261.5 | 0.5205 | 6.1108 | 4.7006 | 0.98 | 137.4633 | 264.2 | 0.5203 | 7.4633 | 5.741 | 157.77 | 97.8
140 | 146.2047 | 270.9 | 0.5397 | 6.2047 | 4.4319 | 0.99 | 147.0444 | 271.5 | 0.5416 | 7.0444 | 5.0317 | 171.32 | 0.89
150 | 156.2404 | 279.7 | 0.5586 | 6.2404 | 4.1603 | 0.98 | 159.1044 | 282.2 | 0.5638 | 9.1044 | 6.0696 | 181.28 | 2.99
160 | 166.9254 | 289.6 | 0.5764 | 6.9254 | 4.3284 | 0.98 | 171.7558 | 293.7 | 0.5848 | 11.7558 | 7.3473 | 191.67 | 5.04
170 | 176.032 | 296.9 | 0.5929 | 6.032 | 3.5482 | 1 | 175.0542 | 296.2 | 0.591 | 5.0542 | 2.9731 | 207.32 | 8.06
180 | 188.774 | 307.7 | 0.6135 | 8.774 | 4.8744 | 1 | 186.1099 | 305.8 | 0.6086 | 6.1099 | 3.3944 | 218.72 | 11.38
190 | 198.3233 | 315.5 | 0.6286 | 8.3233 | 4.3807 | 1 | 197.1284 | 314.6 | 0.6266 | 7.1284 | 3.7518 | 230.41 | 14.69
200 | 201.9619 | 317.9 | 0.6353 | 1.9619 | 0.98093 | 1 | 207.6179 | 326.7 | 0.6355 | 7.6178 | 3.8089 | 241.2 | 17.99

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage