## PURE
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.4907 | 72.3 | 0.1451 | 0.49073 | 4.9073 | 0.67 | 14.3312 | 84.7 | 0.1692 | 4.3312 | 43.3124 | 16.37 | 25.2
20 | 20.0801 | 100.3 | 0.2002 | 0.08006 | 0.4003 | 1 | 20.04 | 100.1 | 0.2002 | 0.04002 | 0.2001 | 29.5 | 42.92
30 | 31.3 | 125 | 0.2504 | 1.3 | 4.3333 | 0.97 | 31.8529 | 126.1 | 0.2526 | 1.8529 | 6.1762 | 38.93 | 52.76
40 | 41.6738 | 144.4 | 0.2886 | 1.6738 | 4.1846 | 1 | 41.5296 | 144.2 | 0.288 | 1.5296 | 3.824 | 52.95 | 60.4
50 | 52.3902 | 162.4 | 0.3226 | 2.3902 | 4.7805 | 0.96 | 54.351 | 164.8 | 0.3298 | 4.351 | 8.7021 | 63.24 | 65.2
60 | 61.5127 | 175.2 | 0.3511 | 1.5127 | 2.5212 | 0.95 | 65.4496 | 180.9 | 0.3618 | 5.4496 | 9.0827 | 73.7 | 69.98
70 | 73.4785 | 191.9 | 0.3829 | 3.4785 | 4.9693 | 1 | 73.1912 | 191.6 | 0.382 | 3.1912 | 4.5589 | 89.2 | 76.25
80 | 82.8244 | 203.7 | 0.4066 | 2.8244 | 3.5305 | 0.97 | 86.154 | 207.7 | 0.4148 | 6.154 | 7.6925 | 98.39 | 79.65
90 | 94.3734 | 217.2 | 0.4345 | 4.3734 | 4.8593 | 0.96 | 98.1687 | 221.8 | 0.4426 | 8.1687 | 9.0763 | 108.55 | 83.27
100 | 101.6779 | 225.6 | 0.4507 | 1.6779 | 1.6779 | 0.97 | 106.7682 | 231.2 | 0.4618 | 6.7682 | 6.7682 | 120.32 | 87.31
110 | 113.0738 | 238 | 0.4751 | 3.0738 | 2.7944 | 0.97 | 118.5108 | 244 | 0.4857 | 8.5108 | 7.7371 | 133.65 | 90.99
120 | 123.2555 | 248.8 | 0.4954 | 3.2555 | 2.7129 | 0.96 | 130.7389 | 256.2 | 0.5103 | 10.7389 | 8.949 | 143.66 | 93.72
130 | 135.0931 | 261.1 | 0.5174 | 5.0931 | 3.9178 | 0.99 | 137.1306 | 262.2 | 0.523 | 7.1306 | 5.4851 | 159.23 | 98.22
140 | 145.7458 | 270.3 | 0.5392 | 5.7458 | 4.1041 | 0.99 | 146.9902 | 271.4 | 0.5416 | 6.9902 | 4.993 | 170.82 | 0.79
150 | 155.1232 | 279.3 | 0.5554 | 5.1232 | 3.4155 | 0.98 | 158.9065 | 282.4 | 0.5627 | 8.9065 | 5.9377 | 179.98 | 2.72
160 | 165.1972 | 287.8 | 0.574 | 5.1972 | 3.2482 | 0.98 | 171.7266 | 293.6 | 0.5849 | 11.7266 | 7.3292 | 191.25 | 4.96
170 | 176.2685 | 297.6 | 0.5923 | 6.2685 | 3.6873 | 1 | 175.5273 | 296.9 | 0.5912 | 5.5273 | 3.2513 | 207.32 | 8.06
180 | 186.3239 | 305.8 | 0.6093 | 6.3239 | 3.5133 | 1 | 185.7429 | 305.8 | 0.6074 | 5.7429 | 3.1905 | 218.72 | 11.38
190 | 198.1351 | 314.7 | 0.6296 | 8.1351 | 4.2816 | 1 | 196.9704 | 314.8 | 0.6257 | 6.9704 | 3.6686 | 229.57 | 14.45
200 | 201.7075 | 317.9 | 0.6345 | 1.7075 | 0.85377 | 1 | 206.8684 | 326.6 | 0.6334 | 6.8684 | 3.4342 | 241.2 | 17.99

## BLEND1
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BLEND2
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## PINPOINT
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
10 | 10.0322 | 71.1 | 0.1411 | 0.03221 | 0.3221 | 0.56 | 18.4794 | 96.7 | 0.1911 | 8.4794 | 84.7937 | 12.3 | 40.81
20 | 20.655 | 102 | 0.2025 | 0.655 | 3.275 | 0.61 | 27.9653 | 118.9 | 0.2352 | 7.9653 | 39.8264 | 20.44 | 58.75
30 | 30.6391 | 124.6 | 0.2459 | 0.63914 | 2.1305 | 0.75 | 39.2975 | 140.7 | 0.2793 | 9.2975 | 30.9917 | 33.03 | 73.59
40 | 41.977 | 145.4 | 0.2887 | 1.977 | 4.9425 | 0.82 | 48.9203 | 157.3 | 0.311 | 8.9203 | 22.3008 | 46.23 | 85.62
50 | 50.1947 | 159.5 | 0.3147 | 0.19465 | 0.3893 | 0.84 | 60.2021 | 174.6 | 0.3448 | 10.2021 | 20.4042 | 58.01 | 94.25
60 | 62.5165 | 176.7 | 0.3538 | 2.5165 | 4.1941 | 0.9 | 68.6162 | 184.8 | 0.3713 | 8.6162 | 14.3604 | 72.66 | 4.01
70 | 70.122 | 188.5 | 0.372 | 0.122 | 0.17429 | 0.89 | 81.1422 | 202.4 | 0.4009 | 11.1422 | 15.9174 | 83.29 | 10.66
80 | 82.6923 | 205.6 | 0.4022 | 2.6923 | 3.3654 | 0.89 | 90.9463 | 214.8 | 0.4234 | 10.9463 | 13.6829 | 94.57 | 17.49
90 | 90.7543 | 214.6 | 0.4229 | 0.75434 | 0.83816 | 0.88 | 105.7022 | 231.6 | 0.4564 | 15.7022 | 17.4469 | 104.49 | 22.22
100 | 104.7901 | 229.3 | 0.457 | 4.7901 | 4.7901 | 0.89 | 115.0484 | 241.8 | 0.4758 | 15.0484 | 15.0484 | 117 | 27.37

## BIPOLAR
Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage