## PURE_1
Calibraiton | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY_2
Calibraiton | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## BIPOLAR_3
Fail
Verify | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
Fail | 10 | 2.9957 | 18.7 | 0.1602 | -7.0043 | -70.0426 | 0 | 3.0751 | 19.1 | 0.161 | -6.9249 | -69.249 | 6.75 | 44.55

## PINPOINT_4
Fail
Verify | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
Pass | 10 | 10.331 | 72.6 | 0.1423 | 0.33098 | 3.3098 | 0.57 | 10.231 | 72.1 | 0.1419 | 0.23099 | 2.3099 | 12.4 | 41.15
Pass | 20 | 20.5897 | 103 | 0.1999 | 0.5897 | 2.9485 | 0.61 | 20.7834 | 103.4 | 0.201 | 0.7834 | 3.917 | 20.33 | 58.58
Pass | 30 | 30.8726 | 127.1 | 0.2429 | 0.87259 | 2.9086 | 0.75 | 30.9709 | 127.4 | 0.2431 | 0.97094 | 3.2365 | 33.05 | 73.62
Pass | 40 | 42.0559 | 147.1 | 0.2859 | 2.0559 | 5.1397 | 0.82 | 42.2596 | 147.4 | 0.2867 | 2.2596 | 5.649 | 46.37 | 85.73
Pass | 50 | 51.0727 | 162.6 | 0.3141 | 1.0727 | 2.1453 | 0.83 | 51.4102 | 163 | 0.3154 | 1.4102 | 2.8204 | 57.68 | 94.02
Pass | 60 | 61.9553 | 178.7 | 0.3467 | 1.9553 | 3.2588 | 0.83 | 61.9542 | 178.8 | 0.3465 | 1.9542 | 3.257 | 67.92 | 0.95
Pass | 70 | 71.8373 | 192.8 | 0.3726 | 1.8373 | 2.6247 | 0.89 | 71.6 | 191.7 | 0.3735 | 1.5999 | 2.2856 | 83.16 | 10.58
Pass | 80 | 83.9449 | 208.3 | 0.403 | 3.9449 | 4.9311 | 0.89 | 83.7055 | 207.5 | 0.4034 | 3.7055 | 4.6319 | 94.32 | 17.34
Pass | 90 | 95.3635 | 222.5 | 0.4286 | 5.3635 | 5.9594 | 0.92 | 95.2247 | 222.8 | 0.4274 | 5.2247 | 5.8052 | 108.55 | 23.95
Pass | 100 | 102.8275 | 230.4 | 0.4463 | 2.8275 | 2.8275 | 0.88 | 102.9902 | 230.3 | 0.4472 | 2.9902 | 2.9902 | 115.22 | 26.66

## PINPOINT_5
Fail
Verify | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
Fail | 10 | 3.2762 | 20.4 | 0.1606 | -6.7238 | -67.2376 | 0 | 102.9902 | 230.3 | 0.4472 | 2.9902 | 2.9902 | 0.00,1.05,1.04,1.02,1.09,1.04 | 44.55
Fail | 20 | 19.9287 | 45.2 | 0.4409 | -0.07132 | -0.3566 | 1.05 | 102.9902 | 230.3 | 0.4472 | 2.9902 | 2.9902 | 39.39 | 14.8
Fail | 30 | 29.5856 | 54.9 | 0.5389 | -0.41439 | -1.3813 | 1.04 | 102.9902 | 230.3 | 0.4472 | 2.9902 | 2.9902 | 55.25 | 39.69
Fail | 40 | 39.5797 | 63.9 | 0.6194 | -0.42034 | -1.0509 | 1.02 | 102.9902 | 230.3 | 0.4472 | 2.9902 | 2.9902 | 70.17 | 57.89
Fail | 50 | 49.8479 | 71.6 | 0.6962 | -0.15208 | -0.30416 | 1.09 | 102.9902 | 230.3 | 0.4472 | 2.9902 | 2.9902 | 91.39 | 80.06
Pass | 60 | 60.217 | 79.4 | 0.7584 | 0.21696 | 0.3616 | 1.04 | 102.9902 | 230.3 | 0.4472 | 2.9902 | 2.9902 | 103.63 | 91.27

## BIPOLAR_6
Verify | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage

## SPRAY_7
Fail
Verify | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
Pass | 10 | 10.5056 | 103.3 | 0.1017 | 0.50561 | 5.0561 | 0.61 | 10.5149 | 103.8 | 0.1013 | 0.51494 | 5.1494 | 13.05 | 41.52
Pass | 20 | 21.3003 | 147 | 0.1449 | 1.3003 | 6.5015 | 0.75 | 21.3148 | 147.1 | 0.1449 | 1.3148 | 6.5739 | 24.77 | 57.3
Pass | 30 | 30.9214 | 177.2 | 0.1745 | 0.9214 | 3.0713 | 0.8 | 31.0095 | 177.4 | 0.1748 | 1.0095 | 3.3651 | 36.89 | 66.67
Pass | 40 | 41.5514 | 205.7 | 0.202 | 1.5514 | 3.8785 | 0.84 | 41.5133 | 204.6 | 0.2029 | 1.5133 | 3.7833 | 49.67 | 73.72
Pass | 50 | 51.597 | 227.5 | 0.2268 | 1.597 | 3.194 | 0.86 | 51.7779 | 228.6 | 0.2265 | 1.7779 | 3.5558 | 61.43 | 81.4
Pass | 60 | 61.9952 | 251.4 | 0.2466 | 1.9952 | 3.3254 | 0.89 | 61.948 | 250.7 | 0.2471 | 1.948 | 3.2466 | 74.67 | 88.76
Pass | 70 | 73.0606 | 271.5 | 0.2691 | 3.0606 | 4.3724 | 0.92 | 73.8177 | 273.5 | 0.2699 | 3.8176 | 5.4538 | 86.67 | 95.51
Pass | 80 | 83.6608 | 295.1 | 0.2835 | 3.6608 | 4.5761 | 0.93 | 81.4387 | 289.2 | 0.2816 | 1.4387 | 1.7984 | 96.42 | 0.92

## SPRAY_8
Fail
Verify | Except Power | Measure Power | Measure Voltage | Measure Current | Power Difference | Power Tolerance | Callibration Data | Measuer Power First | Measuer Voltage First | Measuer Current First | Power Difference First | Power Tolerance First | Mode Watt | Mode Voltage
Fail | 10 | 2.9853 | 18.6 | 0.1605 | -7.0147 | -70.147 | 0 | 81.4387 | 289.2 | 0.2816 | 1.4387 | 1.7984 | 0.00,1.05,1.04,1.02,1.09,1.04 | 44.55
Fail | 20 | 19.6179 | 44.8 | 0.4379 | -0.38208 | -1.9104 | 1.05 | 81.4387 | 289.2 | 0.2816 | 1.4387 | 1.7984 | 39.39 | 14.8